## Slide 1: Gamma & P.A.C Measurements with Nexus-E: A study in compression and stim effects

## Slide 2
Mean PSD during Dyskinetic Part of File
Presence of peak at 70 Hz
Stim On (4V, 140 Hz)
Testing streaming with compression off and compression on

## Slide 3: Baseline PSD without Pxx channels on (streaming from time domain data): No File playing in Saline Tank

## Slide 4: PSD without Pxx channels on (streaming from time domain data): Non-dyskinesia File (no 70 Hz)

## Slide 5: PSD without Pxx channels on (streaming from time domain data): Dyskinesia File (70 Hz)

## Slide 6: Trying different Gamma Configurations
Sense differences between dyskinesia file (red) and non-dyskinesia file (blue)
Different configurations of power channels
Optimize Sensing Channel CF / BW
70 +/- 2.5 Hz
72.5 +/- 2.5 Hz
75 +/- 2.5 Hz
75 +/- 8 Hz

## Slide 7: 1D Gamma Detector
Non Dysk
Dysk
Detection signals of high gamma for different stim levels
Mean Gamma Signals from power channels for different stim levels
Thresh.
No Gamma
Gamma
File played into saline tank to generate above figure (concatenated non-dyskinesia and dyskinesia file together)

## Slide 8: Design 2D detector that uses presence of stim to calculate threshold:
Non Dysk
Dysk
Tracking Stim with (Px2)
~1V during dysk (0-75 sec)
2V during non-dysk
Detect Events Occur When Stim is at ~1V,
Detect Events ALSO Occur When Stim is at ~2V,

## Slide 9: Phase Amplitude Coupling
Can it be sensed w/ Nexus-E?
Use example files
  One w/ high PAC (stim off) , one w/ low PAC (stim on)
Issues :
  Stim artifact bleeding into P.A.C.
  105 Hz artifact bleeding into P.A.C. calculation
  Dropped packets during streaming create P.A.C artifacts (H.F amp. Coupled to L.F. phase)
  BP filters can create edge effects in P.A.C plot
Approaches :
  Band pass filtering
  Removal of sections /w dropped packets
  Stim on different lead than sensing (case of STN stim, cortical sensing)
  Use longer time segment than 400 ms (concatenate data)

## Slide 10: Going back to streaming, 2 leads, being aware of dropped packets and bad BP filter edge effects
Unfiltered
BP filtered: [5 7 70 90], 3dB & 30dB

## Slide 11: Using BP filter and different time segments, does the difference between high and low PAC improve
Notes: Y axis is constant offset to mean PAC, and Y axis is -1*mean(PAC) : 2*mean(PAC)